## Slide 1
MEDICAL DEVICE SECURITY WORKSHOP
Threat Modeling a Medical Device
Hands-on threat modeling of the NeuroScan 3000 — a hybrid cloud + on-premise imaging system
Based on the MITRE Playbook for Threat Modeling Medical Devices

## Slide 2
WHY THREAT MODEL A MEDICAL DEVICE?
The stakes are uniquely high
Patients
A compromised device can directly harm patients — a wrong diagnosis acted on, or care denied in an emergency.
Ransomware
Attacks on hospitals have demonstrably delayed care and put patients at risk.
Data
Medical records are among the most valuable records on the black market.
The law
Since 2023, cybersecurity is a statutory FDA requirement — not optional.

## Slide 3
THE METHOD
Threat modeling answers four questions
1
What are we working on?
Scope the system — assets, entry points, and who might attack it.
2
What can go wrong?
Identify threats as attacker stories, then assess their risk.
3
What are we going to do about it?
Define concrete mitigations and controls.
4
Did we do a good job?
Review completeness, present, and find the gaps.

## Slide 4
NOT JUST IT SECURITY
What makes medical devices different
Aspect
General IT
Medical Devices
Primary risk
Data breach, downtime
Patient harm, misdiagnosis
Availability
Important
Often life-critical
Patching
Weekly / monthly
12–18 months (clearance)
Regulation
GDPR, CCPA
FDA, EU MDR, IEC 81001-5-1
Attackers
Opportunistic
Includes nation-states

## Slide 5
TODAY'S SCENARIO
The NeuroScan 3000
A hybrid MRI-based neurological imaging system. Part sits in the hospital; part runs in the cloud. You will threat model the whole system.
Imaging unit + acquisition workstation (on-prem)
Local DICOM server (on-prem)
Cloud AI diagnostic service
Remote vendor support channel
Hospital EMR — a partially-trusted dependency
HOSPITAL
Imaging unit
Acquisition workstation
Local DICOM server
CLOUD
Cloud AI diagnostic service

## Slide 6
HOW THE TWO HOURS RUN
Run of show
15m
Introduction
Why medical device security — and the scenario
20m
Q1 · What are we working on?
Scope, assets, entry points, actors
35m
Q2 · What can go wrong?
Attacker stories + risk scoring
20m
Q3 · What will we do about it?
Mitigations and trade-offs
20m
Q4 · Did we do a good job?
Review + group presentations
10m
Debrief & wrap-up
What was missed, and what's next

## Slide 7
GROUND RULES
How we'll work
Groups
Work in groups of 3–5. Different views make a better model.
One template
Fill in the shared worksheet as you go through Q1–Q4.
Be specific
Vague threats and vague fixes don't help. Name the component.
AI feedback
Submit your template at the end — an AI evaluator gives written feedback.

## Slide 8
1
What are we working on?
Build a shared map of the system before you hunt for threats.
Agree what's in scope — and what's out (and why).
List the assets worth protecting; mark each C, I, or A.
Map the entry points — every door data or commands come through.
Name the threat actors — insiders, ransomware, nation-states, vendors.

## Slide 9
Q1 REFERENCE · LEAVE THIS UP WHILE YOU WORK
Q1 · Example & what the FDA expects
Worked example — scope
In scope: imaging unit + firmware, acquisition workstation, local DICOM server, cloud AI service, update delivery
Out of scope: hospital EMR (partially trusted); hospital network (threat boundary)
Asset A-01: patient DICOM images (PHI) — Confidentiality + Integrity
Entry point: remote support (VPN + SSH, shared credential) — HIGH RISK
What the FDA expects
Scope the whole medical device system — device + cloud + interoperable parts (§V.A.3)
Capture third-party / COTS components now — they become your SBOM (§524B(b)(3))
Give assets & entry points stable IDs — the anchor for later traceability
Red flag: a boundary drawn around the device only, ignoring the cloud service

## Slide 10
2
What can go wrong?
Think like an attacker. Write ATTACKER STORIES.
As a [bad actor], I want to [do something bad] via [method], so that [goal].
Example: As a rogue vendor technician, I want to change the AI model via the remote support channel, so that the system returns false negatives and patients are misdiagnosed.
Aim for at least 8 stories across different parts of the system.
Tag each story's impact type — S = patient safety · P = privacy/PHI · A = availability.

## Slide 11
Q2 · CONTINUED
Then score each story
Exploitability
how feasible is the attack?
1 Low — rare access / high skill
2 Moderate — remote, known flaw
3 High — trivial, public exploit
×
Severity of harm
safety · privacy · availability
1 Minor — no real harm
2 Significant — delay, breach, downtime
3 Serious — patient harm or major breach
Risk
exploitability × severity
1–2 Low
3–4 Medium
6–9 High
Patient safety rule
If a story could directly harm a patient — wrong diagnosis acted on, device down in an emergency — mark it HIGH regardless of the score, and note why.
In a real engagement, teams layer STRIDE (systematic threat enumeration) and CVSS (standardized scoring, adapted for patient-safety severity) on top of this teaching scale.

## Slide 12
Q2 REFERENCE · LEAVE THIS UP WHILE YOU WORK
Q2 · Example & what the FDA expects
Worked example — a scored threat
T-04 (Elevation of privilege): brute-force the no-MFA admin console on the hospital LAN → full workstation control
Exploitability 3 (no MFA, no rate-limit) × Severity 3 (full compromise) = 9
Patient-safety override: Yes → High
Rationale is about exploitability — how feasible — not the odds of an attack
What the FDA expects
Judge security risk on exploitability, not probability — not the ISO 14971 approach
Keep a severity-of-patient-harm axis → a controlled / uncontrolled call
A patient-safety override can outrank the numeric score
Cybersecurity risk scales independently of software / safety risk

## Slide 13
3
What are we going to do about it?
For each high-priority threat, define at least one concrete mitigation.
Preventive · Detective · Corrective · Compensating — pick the right kind.
Be specific: not "add encryption" but "enable TLS on the DICOM link."
Note the residual risk that remains after the fix.
Flag regulatory impact — could a change need a new FDA submission?

## Slide 14
Q3 REFERENCE · LEAVE THIS UP WHILE YOU WORK
Q3 · Example & what the FDA expects
Worked example — a mitigation
M-01 addresses T-04 & T-07 — type: Preventive
Enforce TOTP MFA on the admin console; replace the shared support credential with per-engineer PKI certs; rate-limit logins
Residual risk: credential theft still possible if the device is already compromised
Regulatory note: software change only — likely no new clearance
What the FDA expects
The risk assessment consumes the threat model; controls are evaluated against threats (§V.A.2)
Record pre- and post-mitigation (residual) risk, with acceptance criteria
Don't let a security control create a new safety risk — the safety ↔ security interface
Document the scoring method, not just the numbers

## Slide 15
4
Did we do a good job?
A threat model only has value if it's honest and complete.
Run the completeness checklist in your template.
Present your top 3 threats, your most surprising find, your hardest trade-off.
Other groups: ask "did you consider…?" — the goal is to find gaps.
The best outcome is another group spotting something you missed.

## Slide 16
Q4 REFERENCE · LEAVE THIS UP WHILE YOU WORK
Q4 · Example & what the FDA expects
Worked example — review & lifecycle
Top-3 risks named; most-surprising = T-09 (rogue insider alters the AI model)
Hardest trade-off: fixing T-01 properly needs secure boot → a new 510(k)
Open question logged: SLA to rotate the shared support credential?
Model kept in version control so every change is diffable
What the FDA expects
Maintain the model across the whole lifecycle — "maintain" is statutory (§524B(b)(2))
Re-trigger on: new feature, new interface, a disclosed vulnerability, an incident
Show end-to-end traceability: threat model ↔ risk assessment ↔ SBOM ↔ testing
Assess per fielded version; track patch % and time-to-patch metrics

## Slide 17
THE REGULATORY REALITY
What the FDA is looking for
Exploitability, not odds
Security risk is judged on how exploitable you are — not on the probability of an attack.
Patient harm first
Severity of patient harm drives the risk, and can override a low score.
A living document
The model must be maintained across the device's whole lifecycle — not filed once.
Traceability
Threat model → risk assessment → SBOM → testing must all connect.

## Slide 18
You are the attacker before the attacker.
Find the problems before a patient gets hurt.
When you're done: name your file threat-model-[team].docx and submit as instructed.